## Valve Information
Human aortic (cellular and decellularised) root details
| Valve ID | Valve Internal Diameter measured with obtratour (mm) | Age (Years) | Sex
| HA044 | 21 | 76 | F
Cellular | HA045 | 22 | 56 | M
| HA057 | 19 | 27 | M
| HA049 | 23 | 50 | M
| HA046 | 23 | 57 | M
Decellularised | HA047 | 21 | 47 | M
| HA048 | 25 | 71 | M
| HA050 | 24 | 75 | F

## Competency testing
Data to support section 4.1.2 in manuscript | Hydrodynamic performance I: competency test under static back pressure for human aortic roots
| Cellular | Mean | Valve | Decellularised | Mean | Valve
| Aortic | Flow Rate | Competenct | Aortic | Flow Rate | Competenet
| Roots | ml/sec | ? | Roots | ml/Sec | ?
| HA044 | <0.85 | * | HA046 | 9.44 | X
| HA045 | <0.85 | * | HA047 | 1.34 | X
| HA057 | <0.85 | * | HA048 | 18.48 | X
| HA049 | 1.78 | X | HA050 | 2.84 | X
* Valve competent (leakage rate ≤ 0.85 ml/sec)
X Valve is not competent (leakage rate > 0.85 ml/sec)
Root Diameters for All Aortic Valves at 120 mmHg Pressure
| Cellular |  | Decellularised
| Valve ID | Root Diameter | Valve ID | Root Diameter
|  | mm |  | mm
| HA045 | 31 | HA048 | 42
| HA044 | 32 | HA046 | 32
| HA049 | 34 | HA047 | 35
| HA057 | 28 | HA050 | 28

## Hydrodynamic
Data to support section 4.1.2 in manuscript
Hydrodynamic performance II: pulsatile flow testing of human aortic roots
|  |  | Mean |  | 95 % Confidence Limits |  | Cellular | Effective Orifice Area EOA(PF) | Decellularised | Effective Orifice Area EOA(PF)
| Aortic Heart Valves | Heart | RMS Flow | Transvalvular Pressure | RMS Flow | Transvalvular Pressure |  | EOA cm2 |  | EOA cm2
|  | Rate | average
|  | bpm | ml/sec | mmHg | ml/sec | mmHg | HA044 | 1.1603 | HA046 | 1.55
|  |  | 0 | 0 | 0 | 0 | HA045 | 1.2208 | HA047 | 1.5163
| Cellular | 60 | 136.94 | 7.729 | 19.88 | 2.5906 | HA057 | 1.1378 | HA048 | 1.99
|  | 72 | 213.307 | 11.62 | 13.64 | 3.8 | HA049 | 1.6577 | HA050 | 1.1613
|  | 80 | 265.759 | 15.63 | 12.819 | 5.2776
[7 empty] | Mean | 1.2942 | Mean | 1.5544
[7 empty] | STDEV | 0.24485 | STDEV | 0.33947
|  |  | 0 | 0 | 0 | 0 | 95 CL | 0.38962 | 95 CL | 0.54017
| Dcellularised | 60 | 147.513 | 5.9574 | 17.8826 | 1.9074
|  | 72 | 205.0512 | 9.3676 | 18.6485 | 3.2921
|  | 80 | 262.8603 | 12.1451 | 17.6435 | 4.6794

## Dilation
Data to support section 4.1.2 in manuscript
Biomechanical performance I: expansion characteristics (dilation testing) of human aortic roots at approximately 40 mmHg pressure
|  | Static Internal Pressure | Dilation |  | Static Internal Pressure | Dilation
| Cellular | mmHg | % | Decellularised | mmHg | %
| HA045 | 36.54 | 16.37 | HA046 | 36.9 | 13.71
| HA049 | 39.5 | 15.92 | HA047 | 40.4 | 11.67
| HA057 | 38.88 | 18.74 | HA050 | 37.6 | 12.35
| HA044 | 38.86 | 10.66
| Average |  | 15.4225 | Average |  | 12.5767
| STDEV |  | 3.4075 | STDEV |  | 1.0387
| 95CL |  | 5.422 | 95CL |  | 2.5803

## Tensile
Data to support section 4.1.2 in manuscript ( Table 2 in Manuscript)
Biomechanical performance II: material properties (uniaxial tensile testing) of human aortic roots
Wall | [5 empty] | Leaflets
|  | Collagen Phase Slope | Elastin Phase Slope | Ultimate Tensile Stress | Failure Strain |  |  |  | Collagen Phase Slope | Elastin Phase Slope | Ultimate Tensile Stress | Failure Strain %
|  | Ec Mpa | Ee Mpa | σUTS Mpa | ɛ failure % |  |  |  | Ec Mpa | Ee Mpa | UTS Mpa | ɛ failure%
Cellular | [6 empty] | Cellular | HA056 | 1.4231 | 0.01514 | 0.233 | 40.55%
Wall | HA056 | 2.0053 | 0.099455 | 1.1033 | 97.83% |  | Leaflet | HA059 | 1.2 | 0.01121 | 0.1086 | 31.66%
Axial | HA059 | 0.83481 | 0.11519 | 0.46866 | 80.83% |  | Radial | HA060 | 0.56292 | 0.029189 | 0.1366 | 29.49%
| HA060 | 0.93173 | 0.082106 | 0.479 | 93.32% |  |  | HA061 | 2.789 | 0.17403 | 0.315 | 18.33%
| HA061 | 4.7926 | 0.44431 | 1.3307 | 46.33% |  |  | HA045 | 1.1002 | 0.048712 | 0.1747 | 26.94%
| HA045 | 1.5852 | 0.075675 | 0.6384 | 85.16% |  |  | HA044 | 0.72251 | 0.021455 | 0.0976 | 25.27%
| HA044 | 0.99818 | 0.12709 | 0.5 | 81.16% |  |  | HA049 | 1.2387 | 0.034338 | 0.1532 | 36.38%
| HA049 | 1.1672 | 0.087515 | 0.60788 | 99.33% |  |  | HA057 | 1.3956 | 0.023015 | 0.28869 | 49.72%
| HA057 | 1.4416 | 0.083622 | 0.6649 | 104.3% |  |  | Mean | 1.304 | 0.044636 | 0.18842 | 32.29%
| Mean | 1.7196 | 0.13937 | 0.7241 | 86.04% |  |  | SD | 0.67313 | 0.053582 | 0.081764 | 9.782%
| SD | 1.3011 | 0.12447 | 0.31885 | 18.25% |  |  | 95CL | 0.56275 | 0.044795 | 0.068356 | 8.178%
| 95CL | 1.0877 | 0.10406 | 0.26657 | 15.26%
Decellularised | HA048 | 2.1117 | 0.1368 | 0.5954 | 55.33% |  | Decellularised | HA048 | 2.1867 | 0.055699 | 0.44281 | 33.61%
Wall | HA046 | 1.4537 | 0.09128 | 0.6169 | 79% |  | Leaflet | HA046 | 3.6548 | 0.092288 | 0.74633 | 36.67%
Axial | HA050 | 2.8746 | 0.32044 | 1.1902 | 60.83% |  | Radial | HA050 | 4.3488 | 0.30812 | 0.51489 | 22.22%
| HA047 | 1.4795 | 0.094048 | 0.8682 | 118% |  |  | HA047 | 6.0667 | 0.13242 | 0.6377 | 26.94%
| Mean | 1.9798 | 0.16064 | 0.81766 | 78.29% |  |  | Mean | 4.0643 | 0.14713 | 0.58543 | 29.86%
| SD | 0.6696 | 0.10855 | 0.2775 | 28.33% |  |  | SD | 1.6107 | 0.11181 | 0.13409 | 6.515%
| 95CL | 1.0655 | 0.17273 | 0.44156 | 45.09% |  |  | 95CL | 2.563 | 0.17791 | 0.21336 | 10.37%
[7 empty] | Cellular | HA056 | 5.8251 | 0.20788 | 0.9098 | 25.99%
Cellular | HA056 | 3.3337 | 0.16286 | 3.387 | 158.2% |  | Leaflet | HA059 | 5.0527 | 0.21085 | 0.52366 | 18.33%
Wall | HA059 | 2.0733 | 0.13343 | 1.3539 | 108.7% |  | Circumferential | HA060 | 6.1128 | 0.24818 | 1.247 | 29.33%
Circumferential | HA060 | 3.2966 | 0.21095 | 1.33 | 90.5% |  |  | HA061 | 3.4651 | 0.089559 | 0.82155 | 35%
| HA061 | 3.2716 | 0.13901 | 0.82155 | 35% |  |  | HA045 | 18.0334 | 0.8971 | 2.1793 | 18.83%
| HA045 | 4.0934 | 0.19071 | 0.909 | 48.66% |  |  | HA044 | 9.9019 | 0.48594 | 1.2203 | 18.99%
| HA044 | 4.7552 | 0.38409 | 1.434 | 52.83% |  |  | HA049 | 20.33 | xx | 2.0079 | 12.49%
| HA049 | 2.5164 | 0.1824 | 0.9783 | 79.66% |  |  | HA057 | 26.5512 | xx | 2.305 | 11.33%
| HA057 | 2.5319 | 0.15329 | 1.5366 | 139.7% |  |  | Mean | 11.909 | 0.35659 | 1.4018 | 21.29%
| Mean | 3.234 | 0.19459 | 1.4688 | 89.14% |  |  | SD | 8.5837 | 0.29506 | 0.67553 | 8.207%
| SD | 0.88018 | 0.080961 | 0.81859 | 44.2% |  |  | 95CL | 7.1762 | 0.24667 | 0.56476 | 6.861%
| 95CL | 0.73585 | 0.067685 | 0.68436 | 36.95%
[7 empty] | Decellularised | HA048 | 16.6419 | XX | 3.1802 | 17.16%
Decellularised | HA048 | 4.0307 | 0.22233 | 1.081 | 51.99% |  | Leaflet | HA046 | 16.6419 | 1.3295 | 1.7894 | 17.83%
Wall | HA046 | 2.732 | 0.16766 | 1.4385 | 106.3% |  | Circumferential | HA050 | 28.7428 | 2.0048 | 2.6749 | 14.66%
Circumferential | HA050 | 6.5908 | 0.28094 | 1.7 | 39.66% |  |  | HA047 | 20.4355 | 1.9335 | 1.99 | 15.99%
| HA047 | 3.6062 | 0.16688 | 1.535 | 96% |  |  | Mean | 20.6155 | 1.7559 | 2.4086 | 16.41%
| Mean | 4.2399 | 0.20945 | 1.4386 | 73.5% |  |  | SD | 5.7057 | 0.37103 | 0.63899 | 1.393%
| SD | 1.6579 | 0.054271 | 0.26173 | 32.62% |  |  | 95CL | 9.0791 | 0.92169 | 1.0168 | 2.216%
| 95CL | 2.6381 | 0.086357 | 0.41646 | 51.9%
XX The initial linear region of the stress - strain graph was used to derive elastin phase modulus and it was not distinct for the specimens. Hence these samples were omitted.

## Suture-pull out
Data to support section 4.1.2 in manuscript
Biomechanical performance II: material properties (suture pull-out testing) for human aortic roots
Aortic
Cellular | Maximum suture pull-out force | Decellularised | Maximum suture pull-out force(N)
Wall | (N) | Wall
HA056 | 9.3 | HA048 | 5.6029
HA059 | 6.27 | HA046 | 6.5248
HA060 | 10.95 | HA047 | 8.3359
Ha061 | 13.23 | HA050 | 6.0736
HA045 | 4.4308
HA044 | 4.6274 | Average | 6.6343
HA049 | 1.8066 | STDEV | 1.1952
HA057 | 4.3311 | 95CL | 1.9019
Average | 6.8682
STDEV | 3.8992
95CL | 3.2598
Cellular | Maximum suture pull-out force | Decellularised | Maximum
Myocardial | (N) | Myocardial | Load (N)
HA056 | 6.33
Ha059 | 5.3 | HA048 | 5.9858
HA060 | 4.54 | HA046 | 2.5294
HA061 | 5.26 | HA047 | 13.1114
HA045 | 4.726 | HA050 | 2.8344
HA044 | 8
HA049 | 7.0449 | Average | 6.1152
HA057 | 4.5155 | STDEV | 4.9189
|  | 95CL | 7.827
Average | 5.7145
STDEV | 1.2842
95CL | 1.0737